|  | THERAPY ACCELERATOR INITIAL COMMISSIONING SUMMARY State Form 46886 (R2 / 5-12) INDIANA STATE DEPARTMENT OF HEALTH MEDICAL RADIOLOGY SERVICES |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Facility registration number |  |  |  | Name of facility |  |  |  |  |  |  |  |  |  | Date of inspection (month, day, year) |  |  |
| Machine number |  | Machine design (use codes) |  |  |  |  | Location |  |  |  |  |  | Manufacturer (use codes) |  |  |  |
| Date of manufacture (month, year) |  | Model number |  |  |  |  |  |  | Serial number |  |  |  |  |  |  |  |
| RADIATION SAFETY SURVEY |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Name of qualified physicist who performed shielding calculations |  |  |  |  | Physicist number |  |  |  |  |  | Shielding document submitted? Yes No |  |  |  |  |  |
| Workload assumptions |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Measurement set-up |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date of radiation safety instrumentation / calibration (month, day, year) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| RESULTS OF RADIATION SURVEY |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Location (Submit a facility layout indicating where the measurements were taken.) |  |  |  |  |  |  | mRem 0o | mRem 90o |  |  |  | mRem 180o |  |  | mRem 270o | mRem per week |
| Console / Control Area |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Primary Barrier (90o Gantry Angle) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Primary Barrier (270o Gantry Angle) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Roof |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Door |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Secondary Barrier |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Secondary Barrier |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Secondary Barrier |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Secondary Barrier |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| HVAC Ductwork |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Accessible Conduits |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| DOSIMETRY SYSTEM AT FACILITY |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Manufacturer of cylindrical chamber |  |  | Model of cylindrical chamber |  |  |  |  |  |  | Date of ADCL calibration (month, day, year) |  |  |  |  |  |  |
| Manufacturer of parallel plate chamber |  |  | Model of parallel plate chamber |  |  |  |  |  |  | Date of N gas derivation (month, day, year) |  |  |  |  |  |  |
| Manufacturer of electrometer |  |  | Model of electrometer |  |  |  |  |  |  | Date of ADCL calibration (month, day, year) |  |  |  |  |  |  |
| Date of aneroid barometer intercomparison (month, day, year) |  |  |  |  |  | Date of initial calibration (month, day, year) |  |  |  |  |  |  |  |  |  |  |
| Calibration protocol |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| QUALITY MANAGEMENT PROGRAM |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Submit the following departmental quality assurance documents: - Treatment planning computer and dose calculated algorithm QA procedures - Patient chart review policy and procedures - Weekly output constancy check policy and procedure - Monthly output spot check procedure |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

* The qualified radiation oncology physicist shall specify tolerance values based upon accelerator manufacture specifications and the most recent published

standards.  Corrective action is required for measured data that exceeds the stated tolerance.

** Energy (bending magnet current), flatness, symmetry, temperature and compensation, et cetera.

*** Target slide or scattering foil, dose chamber slide, dose rate, et cetera.

| FACILITY REQUIREMENTS |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Satisfactory |  |  |  | Unsatisfactory |  | Not Applicable |  |
| Warning Light at Entrance to Accelerator Room Operational |  |  |  |  |  |  |  |  |  |  |
| ACCELERATOR REQUIREMENTS |  |  |  |  |  |  |  |  |  |  |
|  |  |  | Satisfactory |  |  |  | Unsatisfactory |  | Not Applicable |  |
| Absorbed Dose due to Leakage Radiation in the Patient Plane |  |  |  |  |  |  |  |  |  |  |
| Transmission Through Collimators |  |  |  |  |  |  |  |  |  |  |
| Removable Wedges Clearly Identified |  |  |  |  |  |  |  |  |  |  |
| X-ray Contamination of Electrons Beams |  |  |  |  |  |  |  |  |  |  |
| Absorbed Dose at the Surface |  |  |  |  |  |  |  |  |  |  |
| Dose Monitoring System: |  |  |  |  |  |  |  |  |  |  |
| Incorporated into two independent dose monitoring systems |  |  |  |  |  |  |  |  |  |  |
| Monitor units displayed in the event of a power failure |  |  |  |  |  |  |  |  |  |  |
| A symmetry of beam monitored and interlocked |  |  |  |  |  |  |  |  |  |  |
| Secondary dose monitoring system able to terminate irradiation |  |  |  |  |  |  |  |  |  |  |
| Timer terminates irradiation if dose monitoring systems fail |  |  |  |  |  |  |  |  |  |  |
| Selection of Radiation Mode |  |  |  |  |  |  |  |  |  |  |
| Selection of Energy |  |  |  |  |  |  |  |  |  |  |
| Selection of Fixed or Rotational Mode |  |  |  |  |  |  |  |  |  |  |
| MECHANICAL PARAMETERS |  |  |  |  |  |  |  |  |  |  |
|  |  | Satisfactory |  | Unsatisfactory |  |  | Tolerance * |  |  |  |
| Gantry Rotation Isocenter |  |  |  |  |  |  |  |  |  |  |
| Collimator Rotation Isocenter |  |  |  |  |  |  |  |  |  |  |
| Table (Couch) Rotation Isocenter |  |  |  |  |  |  |  |  |  |  |
| Field Size Readouts |  |  |  |  |  |  |  |  |  |  |
| Gantry Angle Readouts |  |  |  |  |  |  |  |  |  |  |
| Collimator Angle Indicator |  |  |  |  |  |  |  |  |  |  |
| Optical Distance Indicator (ODI) Accuracy and Linearity |  |  |  |  |  |  |  |  |  |  |
| Light / Radiation Field Congruency |  |  |  |  |  |  |  |  |  |  |
| Laser Alignment System Accuracy |  |  |  |  |  |  |  |  |  |  |
| Emergency Off Switches |  | Operational? |  |  | Yes |  |  | No |  |  |
| Door Interlock |  | Operational? |  |  | Yes |  |  | No |  |  |
| Dosimetry Interlocks ** |  | Operational? |  |  | Yes |  |  | No |  |  |
| Safety Interlocks *** |  | Operational? |  |  | Yes |  |  | No |  |  |
| Audio Communications System |  | Operational? |  |  | Yes |  |  | No |  |  |
| Visual Monitor System |  | Operational? |  |  | Yes |  |  | No |  |  |
| PHOTON BEAM PARAMETERS |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  | Tolerance * |
| Nominal Energy (MV) |  |  |  |  |  |  |  |  |  |  |
| Ionization Ratio |  |  |  |  |  |  |  |  |  |  |
| Flatness | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| Symmetry | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| Field Size Factors | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| PDD / TPR Values | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| Output / MU Reproducibility | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| Monitor Chamber Linearity | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| Wedge / Compensator Transmission Factors | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| Tray / Custom Block Transmission Factors | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| Open Field Isodose Lines (central axis) | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| Wedge Isodose Lines (central axis) | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |  | Satisfactory Unsatisfactory |  |  |  |  |
| Output Calibration | Yes No | Yes No |  |  |  | Yes No |  |  |  | Performed |
| Constancy Check Device Calibration | Yes No | Yes No |  |  |  | Yes No |  |  |  | Performed |

* The qualified radiation oncology physicist shall specify tolerance values based upon accelerator manufacture specifications and the most recent published

standards.  Corrective action is required for measured data that exceeds the stated tolerance.

| ELECTRON BEAM PARAMETERS |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | Tolerance * |
| Nominal Energy (MeV) |  |  |  |  |  |
| Mean Incident Energy |  |  |  |  |  |
| Flatness (Cone Size ) | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |
| Symmetry | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |
| Output / MU Reproducibility | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |
| Output Calibration | Yes No | Yes No | Yes No |  | Performed |
| Constancy Check Device Calibration | Yes No | Yes No | Yes No |  | Performed |
| ELECTRON BEAM PARAMETERS |  |  |  |  |  |
|  |  |  |  |  | Tolerance * |
| Nominal Energy (MeV) |  |  |  |  |  |
| Mean Incident Energy |  |  |  |  |  |
| Flatness (Cone Size ) | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |
| Symmetry | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |
| Output / MU Reproducibility | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |
| Output Calibration | Yes No | Yes No | Yes No |  | Performed |
| Constancy Check Device Calibration | Yes No | Yes No | Yes No |  | Performed |
| ELECTRON BEAM PARAMETERS |  |  |  |  |  |
|  |  |  |  |  | Tolerance * |
| Nominal Energy (MeV) |  |  |  |  |  |
| Mean Incident Energy |  |  |  |  |  |
| Flatness (Cone Size ) | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |
| Symmetry | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |
| Output / MU Reproducibility | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory | Satisfactory Unsatisfactory |  |  |
| Output Calibration | Yes No | Yes No | Yes No |  | Performed |
| Constancy Check Device Calibration | Yes No | Yes No | Yes No |  | Performed |
| Signature of physicist |  |  |  | Date (month, day, year) |  |
| Printed name of physicist |  |  |  | Physicist number |  |